## Slide 1
The Lego® Spectrophotometer
Practical Investigations of Dye
December 2022

## Slide 2
Workshop objectives:
To build a Lego® Spectrophotometer
To gain an understanding of how spectrophotometers work.
To apply simple circuitry to build a functional spectrophotometer.
To plot a calibration curve for Brilliant Blue FCF dye, and determine the concentration of an unknown solution.

## Slide 3
What is Spectrophotometry?
Spectrophotometry is a type of electromagnetic spectroscopy which allows determination of the absorptivity or transmittivity of a substance under investigation. Spectrophotometry is commonly used in:

## Slide 4
How Do Spectrophotometers Work?
1. Polychromatic light source.
2. Monochromator refocuses only one wavelength onto the cuvette.
3. Focussed light passes through the sample with incident intensity, I0. Sample absorbs some light and transmits the rest, with transmission intensity, I. Since some light is absorbed by the sample, I0 > I.
4. Light with transmitted intensity, I, reaches the detector which displays an output of voltage.
5.The resultant voltage can be used to deduce the transmitted intensity. Voltage, V, is directly proportional to transmitted intensity, I.

## Slide 5
Why are objects coloured?
Notes: Most objects will absorb some wavelengths of light better than others and the colour of a substance depends on the wavelengths of light that it absorbs. For a solution to appear as a particular colour, the dissolved compound must absorb the complementary colour to which it appears. Complementary colours lie opposite on a colour wheel.
The interaction of light with the sample is an important consideration in the Lego® spectrophotometer. We must choose a light source that the compound will absorb in order to record accurate optical measurements. To investigate different coloured dyes you need to choose a correctly coloured LED light source. In the case of Brilliant Blue FCF, which appears blue, what colour light is needed?
A: Orange – it is opposite blue on the colour wheel.

## Slide 6
Why do dye molecules absorb light?
Highest occupied molecular orbital
(HOMO)
Lowest unoccupied molecular orbital
(LUMO)
Notes: Electrons in compounds can be excited from the HOMO (highest occupied molecular orbital) to the LUMO (lowest unoccupied molecular orbital) by absorbing energy in the form of a photon (see Figure 4). In the case of many conjugated systems, the photon (packet of energy) absorbed to excite an electron lies in the visible light region, and so the compound appears coloured. The wavelength of the absorbed photon will correspond to the complementary colour of the compound.

## Slide 7
Determination of Brilliant Blue FCF in a sports drink
In this practical, you will determine the concentration of Brilliant Blue FCF dye in a drink of unknown concentration.

## Slide 8
Beer-Lambert Law
The Beer Law states that:
Where
A = sample absorbance
= concentration (normally in mol dm-3),
= path length (1 cm for standard cuvettes)
e = absorption coefficient (normally in mol-1 dm3 cm-1).

## Slide 9
Build the Lego® Spectrophotometer
Important!
1. Always use the resistor.
2. Before anything else: make sure LEDs are working and check which colour is which.
3. Once built, move immediately on to the experimental stage.

## Slide 10
Any questions

## Slide 11
Find out more about the work of the Department of Chemistry at
www.chem.ox.ac.uk
outreach@chem.ox.ac.uk
@ChemOutreachOxf
vector created by starline